## PM2.5
Peak 1 - ug/m3 | [7 empty] | Accuracy | Grade B
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | 0-150 | ±5 && 15%
2021-10-25 04:30 | 16 | 16.5872 | 17.273 | 16.656 | 16.526 | 15.677 | 16.804 | 150-500 | ±5 && 20%
Percent Difference to Avg |  |  | 4.051% | 0.4139% | -0.3696% | -5.642% | 1.299%
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS
Percent Difference to Reference |  |  | 7.652% | 4.018% | 3.234% | -2.039% | 4.902%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within
Peak 2 - ug/m3 | [7 empty] | PM2.5 | No.1 | No.2 | No.3 | No.4 | No.5
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | Avg - Trend | PASS | PASS | PASS | PASS | PASS
2021-11-04 04:30 | 24 | 24.6224 | 25.291 | 24.764 | 24.529 | 23.719 | 24.809 | Avg – Within Range | PASS | PASS | PASS | PASS | PASS
Percent Difference to Avg |  |  | 2.679% | 0.5734% | -0.3801% | -3.738% | 0.755% | Reference - Trend | PASS | PASS | PASS | PASS | PASS
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS | Reference – Within/Outside Scope | Within | Within | Within | Within | Within
Percent Difference to Reference |  |  | 5.238% | 3.133% | 2.18% | -1.178% | 3.315%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within
Peak 3 - ug/m3
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5
2021-11-06 08:00 | 50 | 50.5792 | 51.22 | 50.657 | 50.517 | 49.697 | 50.805
Percent Difference to Avg |  |  | 1.259% | 0.1537% | -0.1231% | -1.76% | 0.4454%
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS
Percent Difference to Reference |  |  | 2.411% | 1.305% | 1.029% | -0.6078% | 1.597%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within

## TVOC
Peak 1 - ppb | [7 empty] | Accuracy (ug/m3) | Grade B
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | 150-600 | ±20 && 15%
2021-10-25 14:00 | 570 | 588.67 | 614.574 | 529.756 | 596.804 | 608.097 | 594.119 | 600-2000 | ±20 && 20%
Percent Difference to Avg |  |  | 4.306% | -10.54% | 1.372% | 3.247% | 0.9214%
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS | Accuracy (ppb) | Grade B
Percent Difference to Reference |  |  | 7.526% | -7.319% | 4.594% | 6.468% | 4.144% | 65-260 | ±8.7 && 15%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within | 440 - 2180 | ±8.7 && 20%
Peak 2 - ppb | [7 empty] | TVOC | No.1 | No.2 | No.3 | No.4 | No.5
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | Avg - Trend | PASS | PASS | PASS | PASS | PASS
2021-10-31 11:30 | 520 | 541.6324 | 558.992 | 487.966 | 557.65 | 548.761 | 554.793 | Avg – Within Range | PASS | PASS | PASS | PASS | PASS
Percent Difference to Avg |  |  | 3.155% | -10.42% | 2.914% | 1.308% | 2.401% | Reference - Trend | PASS | PASS | PASS | PASS | PASS
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS | Reference – Within/Outside Scope | Within | Within | Within | Within | Within
Percent Difference to Reference |  |  | 7.227% | -6.356% | 6.987% | 5.382% | 6.474%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within
Peak 3 - ppb
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5
2021-11-07 11:00 | 500 | 515.616 | 527.208 | 464.316 | 523.978 | 535.729 | 526.849
Percent Difference to Avg |  |  | 2.223% | -10.47% | 1.609% | 3.826% | 2.155%
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS
Percent Difference to Reference |  |  | 5.297% | -7.401% | 4.683% | 6.899% | 5.229%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within

## CO2
Peak 1 - ppm | [7 empty] | Accuracy | Grade B
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | 400-2000 | ±50 && 3%
2021-10-24 11:00 | 1613 | 1644.076 | 1639.45 | 1645.31 | 1640.5 | 1655.49 | 1639.63 | 2000-5000 | ±50 && 5%
Percent Difference to Avg |  |  | -0.2818% | 0.07503% | -0.2177% | 0.6918% | -0.2708%
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS
Percent Difference to Reference |  |  | 1.626% | 1.983% | 1.69% | 2.6% | 1.637%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within
Peak 2 - ppm | [7 empty] | CO2 | No.1 | No.2 | No.3 | No.4 | No.5
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | Avg - Trend | PASS | PASS | PASS | PASS | PASS
2021-10-31 09:30 | 1867 | 1899.62 | 1902.41 | 1908.16 | 1895.36 | 1900.58 | 1891.59 | Avg – Within Range | PASS | PASS | PASS | PASS | PASS
Percent Difference to Avg |  |  | 0.1468% | 0.4486% | -0.2245% | 0.05052% | -0.4236% | Reference - Trend | PASS | PASS | PASS | PASS | PASS
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS | Reference – Within/Outside Scope | Within | Within | Within | Within | Within
Percent Difference to Reference |  |  | 1.879% | 2.181% | 1.508% | 1.783% | 1.308%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within
Peak 3 - ppm
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5
2021-10-28 10:30 | 1355 | 1383.9514 | 1378.641 | 1384.398 | 1387.139 | 1391.25 | 1378.329
Percent Difference to Avg |  |  | -0.3845% | 0.03226% | 0.2301% | 0.526% | -0.4071%
Avg Pass/Fail |  |  | PASS | PASS | PASS | PASS | PASS
Percent Difference to Reference |  |  | 1.73% | 2.146% | 2.344% | 2.64% | 1.707%
Reference Within/Outside Scope |  |  | Within | Within | Within | Within | Within

## Temp
Peak 1 - °C | [8 empty] | Grade B
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | Accuracy | ±1
2021-10-21 01:00 | 25.7 | 27.6794 | 27.768 | 27.329 | 27.885 | 27.679 | 27.736
Difference to Avg |  |  | 0.0886 | -0.3504 | 0.2056 | -0.0004 | 0.0566
Difference to Reference |  |  | 2.068 | 1.629 | 2.185 | 1.979 | 2.036
Peak 2 -°C | [7 empty] | Temperature | No.1 | No.2 | No.3 | No.4 | No.5
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | Avg - Trend | PASS | PASS | PASS | PASS | PASS
2021-10-29 07:30 | 27.7 | 29.2634 | 29.371 | 28.88 | 29.462 | 29.275 | 29.329 | Avg – Within Range | PASS | PASS | PASS | PASS | PASS
Difference to Avg |  |  | 0.1076 | -0.3834 | 0.1986 | 0.0116 | 0.0656 | Reference - Trend | PASS | PASS | PASS | PASS | PASS
Difference to Reference |  |  | 1.671 | 1.18 | 1.762 | 1.575 | 1.629 | Reference – Within/Outside Scope | Outside | Outside | Outside | Outside | Outside
Peak 3 -°C
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5
2021-11-03 01:00 | 26.7 | 28.252 | 28.387 | 27.888 | 28.442 | 28.228 | 28.315
Difference to Avg |  |  | 0.135 | -0.364 | 0.19 | -0.024 | 0.063
Difference to Reference |  |  | 1.687 | 1.188 | 1.742 | 1.528 | 1.615

## Hum
Peak 1 - %RH | [8 empty] | Grade B
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | Accuracy | ±8%
2021-10-25 23:30 | 66.2 | 64.3058 | 64.449 | 64.016 | 64.496 | 64.014 | 64.554
Percent Difference to Avg |  |  | 0.2224% | -0.4517% | 0.2953% | -0.4548% | 0.3852%
Percent Difference to Reference |  |  | -2.68% | -3.354% | -2.608% | -3.358% | -2.518%
Peak 2 - %RH | [7 empty] | Relative Humidity | No.1 | No.2 | No.3 | No.4 | No.5
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5 | Avg - Trend | PASS | PASS | PASS | PASS | PASS
2021-10-31 01:00 | 65.68 | 63.7716 | 63.889 | 63.54 | 63.874 | 63.531 | 64.024 | Avg – Within Range | PASS | PASS | PASS | PASS | PASS
Percent Difference to Avg |  |  | 0.1839% | -0.3638% | 0.1604% | -0.378% | 0.395% | Reference - Trend | PASS | PASS | PASS | PASS | PASS
Percent Difference to Reference |  |  | -2.765% | -3.312% | -2.788% | -3.326% | -2.554% | Reference – Within/Outside Scope | Within | Within | Within | Within | Within
Peak 3 - %RH
Time | Reference | Avg | No.1 | No.2 | No.3 | No.4 | No.5
2021-11-03 00:30 | 66.2 | 64.2992 | 64.459 | 64.033 | 64.357 | 64.151 | 64.496
Percent Difference to Avg |  |  | 0.2482% | -0.4149% | 0.08985% | -0.2308% | 0.3056%
Percent Difference to Reference |  |  | -2.665% | -3.328% | -2.823% | -3.144% | -2.608%